Equipment Required:

1 x high quality pH/mV Meter			4 x 200ml beakers

100ml pH 7.00 buffer solution			3 x Teflon stirrer paddles

100ml pH 4.00 buffer solution			1 x magnetic stirrer

100ml pH 5.00 buffer solution			1 x cantilever electrode arm

1 x calibrated glass thermometer                                1 x fast flow wash bottle (DI water)

1 x temperature controlled water bath (required if the sample value is different from

ambient)

1 x combination pH electrode

Method

- Assemble all equipment.

- Lower fill hole sleeve on the electrode (if fitted) and thoroughly rinse the electrode tip.

- Lower electrode into gently stirred pH 7.00 buffer and allow to stabilise.

- Check the temperature of the calibration solutions and adjust the default reading on the pH meter, if applicable.

- After 1-2 minutes adjust the calibration control on the pH meter to the appropriate pH value. (consult pH buffer temperature versus pH chart).

- Raise electrode from beaker and thoroughly rinse with DI water.

- Lower electrode into gently stirred pH 4.00 buffer and allow to stabilise.

- After 1-2 minutes adjust the slope control on the pH meter to the appropriate temperature corrected value.

NOTE:  	Many modern microprocessor pH meters have automatic buffer recognition. Consult the instrument manual for specific adjustment information.

- Rinse the electrode and repeat stages c) - h) to confirm calibration.

- Rinse the electrode and lower into pH 5.00 buffer.

- After stabilising, record the reading in pH 5.00 buffer.

- Between measurements in the sample, rinse and lower the electrode into the control buffer for comparison with the recorded reading. (remember to check temperature pH versus pH values).