Operating Instructions

Barium Ion Selective Electrode

This Document includes operating instructions for two types of Ion Selective Electrode. A single sensing Half Cell Barium Electrode which requires a separate Reference Electrode, and a Combination Barium Electrode which has the Reference built into the outer shaft.

For extremely low level work the Half Cell plus Reference Electrode should be used. For general applications the Combination is more convenient.

Depending on the type of ISE chosen you will need to disregard some information for instance if you have the Combination Electrode then disregard any comments about reference electrodes.

A quick guide is included with a more involved section following.

Quick Guide to Barium

ISE

The combination ion selective electrode has a solid state PVC polymer membrane with an integral driTEK reference electrode. The electrode is designed for the detection and analysis of Barium ions in aqueous solutions and is suitable for use in the field in the laboratory and in on line analyzers.

This ISE is ideal for the analysis of Sulphate concentrations as an endpoint indicator when using the titrimetric method.

## Installation

Connect the ISE to the mV or Ion meter.

Remove the black protective cap and keep it in a safe place.

The ISE can be used immediately but pre soaking for 5 minutes in a 100ppm Barium solution is recommended.

The ionic strength of the standards and solutions should be kept constant between all standards and samples. This is achieved by the simple addition of an Ionic strength adjustment buffer (ISAB). Ammonium Nitrate is ideal. A typical addition would be 1 ml of 1 molar ISAB to 100 ml of standard and sample.

For low level measurements below around 50ppm in relatively pure samples no ISAB is needed.

No temperature correction is necessary however standards and samples should be measured at the same temperature.

Begin calibration from the lowest concentration standard to avoid cross contamination. Calibration should cover the anticipated range of the samples.

Rinse tip with de-ionised water between measurements.

Avoid strongly acidic or alkaline samples, strong detergents and organic solvents.

## Storage and Maintenance

Avoid touching the membrane surface.

After use rinse with de-ionised water, replace protective cap and store dry in its box.

If performance becomes sluggish, rinse with dilute detergent, rinse with de-ionised water and immerse the tip in a 1000ppm Barium solution for 1 hour.

Specification

Barium is a Divalent Cation.

Refer to the table below for the full specification:

| Parameter | Specification |
| --- | --- |
| Overall length Body Diameter Cap Diameter Connector Cable length Resistance at 25 Deg C Concentration Range Slope Potential Drift Operating pH range Temperature range Endpoint time Interferences. Ions with coefficients above 0.001. | 155 mm 12 mm 16mm BNC 180 mm < 2.5 Meg Ohm 0.6 to 5,600 ppm 24 to 29 mV per decade 2 mV per day 3 to 10 5 to 50 Deg C Typically 20 to 50 seconds Strontium Potassium Sodium Magnesium Calcium Lithium Ammonium |

#### Recommendations for successful analysis

The combination ISE’s can be used with any pH/mV meter or Ion meter. If the meter does not have a BNC socket and you have a BNC electrode please contact you’re your distributor who will arrange to have the correct plug fitted. Adapters are also available if the same electrode has to be used on more than one meter.

Meters with a 0.1 mV resolution are recommended whilst dedicated Ion meters will provide direct concentration readouts saving time and effort in constructing calibration curves and performing calculations. Your distributor can advise on the most suitable meter.

Magnetic stirrer/stirrer bars are recommended for laboratory analysis. Please operate at the lowest constant speed available.

Semi-logarithmic (4-cycle) graph paper is required for preparing calibration curves when you are using a mV meter.

#### Required Solutions

Distilled or de-ionised water will be required to prepare Standards, ISABs and to rinse the electrode between measurements.

1000ppm Stock Standard solution. Used for preparation of Standards. (Prepared by customer)

ISAB. Used to adjust the Ionic strength of all standards and samples. Typical addition is 1 ml of 4 Molar ISAB to 50ml of all standards and samples.

#### Operation

- Connect the electrode to the meter being used for analysis

- Prepare a series (at least 2) of standards that bracket the expected sample concentration. This is best done by serial dilution of the stock solution. Ideally standards should be a decade in concentration apart e.g. 1, 10, and 100ppm.

- Dispense 50 ml of each standard into analytically clean beakers (100 to 150 ml size is perfect)

- Add ISAB/TISAB in the appropriate ratio. As a guide with sample concentrations in the 1 to 1000ppm range 1ml of a 4 Molar ISAB to 50 ml sample is satisfactory. For TISAB(Fluoride analysis only) please read the label.

- Rinse the electrode with de-ionised water and blot dry with a lint free cloth and place in the lowest standard. When the reading is stable record the mV value.

- Repeat step 6 for all subsequent standards proceeding from lowest to highest.

- Plot a calibration curve on semi log paper using mV values on the linear Axis and concentration on the log scale.

- Rinse the electrode in de-ionised water and blot dry. Place the electrode in the sample and record the stable mV value.

- Using the calibration curve determine the unknown sample concentration.

Hints and tips

- Ensure that the temperature of all standards and samples are the same to reduce errors.

- Using a magnetic stirrer for laboratory analysis is recommended but not essential. It is however important to have the stirrer set on a low constant speed which must be reproducible for all measurements.

- Prior to sample measurement ensure that the electrode is thoroughly rinsed with de-ionised water. It is worth performing this rinse twice given the possibility of carryover being greatest in high concentration solutions.

- Prepare standards by serial dilution.

- Make sure your electrode is conditioned by leaving the tip in the lowest concentration standard for 1 hour prior to analysis.

Methods of Analysis

We have described direct potentiometry above.  This method is simplified by using a direct reading ion meter. There are several other methods, which are useful.

Known Addition: An incremental technique where the potential of the sample solution is measured followed by addition of a small volume of a higher concentration standard solution. The new potential is measured and from difference in the two values, and using the known electrode slope, the unknown concentration is determined,

This method is ideal for samples whose matrix is not entirely clean or aqueous. In these

instances comparisons with clean standards is not appropriate thereby making direct potentiometry unsuitable. Known addition works because both standard and sample are measured in the same matrix.

Typical sample volume is 50 ml, typical standard volume is 5 ml. The standard should be approximately 100 time the sample concentration for accurate analysis.

Sample Addition: An incremental technique where the potential of a dilute standard solution is measured followed by the addition of a small volume of more concentrated sample. The new potential is recorded and the difference noted. Using this value (and the predetermined electrode slope) the unknown concentration is determined.

This method is ideal for dirty or viscous samples with an awkward matrix. The sample however needs to be relatively concentrated i.e. at least 100 times the Electrodes linear detection limit. The analysis does have the benefit of only requiring a small volume.

The sample matrix is basically broken down by dilution with the standard and therefore analysis is carried out in the same media.

End Point Titration: Flow Plus combination ISE’s are ideal end point indicators and will produce a significant potential change at the equivalence point. The Ion in question must be contained in the titrand or the titrant and must therefore be in excess or absence at the end point.

Operating Instructions for Barium

Ion Selective Electrode

- Introduction

  - Construction

The electrode consists of an inert fluorocarbon body with a solid state PVC membrane unit. The sensitive membrane consists of liquid ion exchange material immobilised in a poly-(vinyl chloride) matrix. Connection to a millivoltmeter is via a coaxial cable and BNC plug. As an alternative other plugs may be specified.

  - Operating Principles

The electrical potential of an ion selective electrode is a function of the activity of certain ions in an aqueous solution. This potential can only be measured against a reference electrode, such as a saturated calomel electrode (SCE), placed in the same solution. Measurements are made with a high impedance millivoltmeter or expanded scale pH meter. A meter with a digital display is obviously more accurate and most convenient. The barium electrode cannot be used in non-aqueous solvents but the presence of up to 40% organic solvents in an aqueous solution can be tolerated. For further discussion on the principles of ISE measurements, the user is referred to the Technical Bulletin, which is available in demand.

  - Sensitivity

The electrode will respond to uncomplexed barium ion activity over the range 100

10-6M. The linear detection limit is around 5 x 10-5M but careful calibration allows the electrode to be used down to 10-6M. Curvature of the response is observed at lower concentrations because the solubility of the liquid ion-exchanger, leached from the membrane, contributes a small but significant amount of barium to the solution being measured. Figure 1 shows a typical calibration plot of the electrode.

Once the ionic strength is calculated and the activity coefficient found from Table 1, then the activity of the ion of interest is found from :

Activity (Ai) = Concentration (Ci) x Activity Coefficient (f)

  - Stability

The long-term drift in the E0 value of the barium electrode is usually in the region of 1-2mV/day.

  - Selectivity

The major interferent to the selectivity of this electrode is the lead ion, to which, under certain circumstances, the electrode can appear even more selective than for barium. However, its response is non-linear. The figure for selectivity coefficients given here have been determined at barium concentration of 10-3M. The ratios are concentration dependent and these figures should be taken as a guide only.

Interferent					Selectivity Ratio KBa/X

Pb2+							2.4

Ca2+							0.025

Na+							0.0004

K+							0.009

Since barium and lead are unlikely to be found together and the selectivity of the electrode over other commonly occurring ions is sufficiently high, the electrode should prove a useful monitoring probe in real situations.

The electrode should be used in the pH range 2-11.

  - Response Time.

The electrode will respond to a step change in the barium ion activity in solution less than ten seconds. When, however, the electrode is removed from one solution and placed in another of a different barium concentration, it may take up to 30 seconds to achieve a new equilibrium.

The presence of some interfering ions, especially Mg2+ and Ca2+ can cause some increase in the response time. After exposure to strong concentrations of such ions the electrode should be re-conditioned in 10-2M barium solution for about an hour.

- Operating and Handling

  - Initial Filling and Setting Up.

On receipt of the electrode remove it from its packaging and inspect for damage, particularly to the delicate membrane, which forms the sensing end of the electrode. The electrode is now ready for use.

  - Making a Measurement

The electrode should first be calibrated as follows. Clamp the electrode in a vertical position such that the membrane is immersed in a beaker of 10-1M barium solution placed on a magnetic stirrer. Ensure that the liquid does not come above the membrane unit/electrode body junction. Clamp a reference electrode (a single junction Ag/AgCl or SCE is quite adequate) so that its end is immersed in the same solution and connect both electrodes to a pH or a millivolt meter. Record the potential when it becomes steady. Change the test solution sequentially for 10-2M,

10-3M and 10-4M barium solutions and plot the calibration curve, washing and drying the electrode between the solutions.

Unknown solutions can now be measured directly. For further details of methods of measurement, the user is referred to Technical Bulletin E1 available form the electrode supplier.

  - Storage

If the electrode is to be used within a week or so, it can be merely removed from the test solution, washed and dried and be clamped in a vertical position in the air. There is, in principle, no limit to the storage time, but gradual oxidation of the membrane will, however, cause it to harden and lose its response. Shelf life, in its dry state, is better than 12 months.

  - Flow Through Configuration

A conversion kit is available to enable continuous flow-through measurements to be made. The reference contact can be made at a point in the solution stream joined to the flow cell by a continuous stream of liquid usually downstream of the electrode. The flexible nature of the membrane causes the pump pulses to produce rather more 'noise' on the electrode signal than normally found (+0.8mV as opposed to +0.2mV). This can be mitigated by either damping the recorder-input signal or using a pulseless pump.

Values for the slope of this plot should lie in the range 24-29 mV/decade. If the total ionic strength of all solutions is maintained constant by the addition of a concentrated (5 x 10-2M) inert electrolyte, such as Lithium Chloride, the electrode can be calibrated directly in concentration units.

Daily re-calibration is recommended for analytical use.

The activity coefficient of barium ion depends upon the ionic strength of the solution. The following table can be used to calculate the approximate activity coefficients at various ionic strengths.

Ionic Strength		0.001	0.005	0.01	0.05	0.1	0.2

Activity Coeff.	0.905	0.809	0.749	0.57	0.485	0.405

The ionic strength of the solution to be tested can be calculated from the formula:

Ionic Strength		=	½  Σ CiZi2

Where Ci is the concentration of species i and Zi is its charge, and all species, cations and anions, must be taken into account.

- Applications

It is possible to predict that the single greatest application of this electrode will be as the potentiometric sensor in the titration of sulphate ions with barium chloride. Figure 2 shows the resultant titration curve, from which the concentration of sulphate can easily be determined. A potentiometric titration can be performed in turbid coloured or other "difficult" samples directly and the use of an automatic titrator can result in a rapid, precise method for an analysis that has long proved difficult.

The presence of phosphate can be overcome by adding an excess of nitric acid to prevent Ba3 (PO4)2 precipitation.

Reference: A. M. Y. Jaber, G. J. Moody and J. D. R. Thomas; Analyst, 101 (1976) 179-186.

# Text boxes

150mV

300mV

Mono/Reference

Combination

SENTEK

Unit 6 & 7 Crittall Court, Crittall Drive, Springwood Industrial

Estate, Braintree, Essex, CM7 2SE

Tel: + 44 (0) 1376 340 456  Fax:  +44 (0) 1376 340 453  Email: sentekuk@btconnect.com

SENTEK

Unit 6 & 7 Crittall Court, Crittall Drive, Springwood Industrial

Estate, Braintree, Essex, CM7 2SE

Tel: +44 (0) 1376 340 456  Fax: +44 (0) 1376 340 453  Email:  sentekuk@btconnect.com