Operating Instructions

Iodide Ion Selective Electrode

This Document includes operating instructions for two types of Ion Selective Electrode. A single sensing Half Cell Iodide Electrode which requires a separate Reference Electrode, and a Combination Iodide Electrode which has the Reference built into the outer shaft.

For extremely low level work the Half Cell plus Reference Electrode should be used. For general applications the Combination is more convenient.

Depending on the type of ISE chosen you will need to disregard some information for instance if you have the Combination Electrode then disregard any comments about reference electrodes.

A quick guide is included with a more involved section following.

Quick Guide for Iodide

ISE

The combination ion selective electrode has a solid state crysytalline membrane with an integral driTEK reference electrode. The electrode is designed for the detection and analysis of Iodide ions in aqueous solutions and is suitable for use in the field in the laboratory and in on line analysers.

## Installation

Connect the ISE to the mV or ion meter.

Remove the black protective cap and keep it in a safe place.

The ISE can be used immediately but pre soaking for 5 minutes in a 100ppm Iodide solution is recommended.

The ionic strength of the standards and solutions should be kept constant between all standards and samples. This is achieved by the simple addition of an Ionic strength adjustment buffer (ISAB). Potassium Nitrate is ideal. A typical addition would be 1 ml of 1 molar ISAB to 100 ml of standard and sample.

For low level measurements below around 50ppm in relatively pure samples no ISAB is needed.

No temperature correction is necessary however standards and samples should be measured at the same temperature.

Begin calibration from the lowest concentration standard to avoid cross contamination. Calibration should cover the anticipated range of the samples.

Rinse tip with de-ionised water between measurements.

Avoid strongly acidic or alkaline samples, strong detergents and organic solvents.

## Storage and Maintenance

After use rinse with de-ionised water, wipe clean with a tissue or lint free cloth, replace protective cap and store dry in its box.

If performance becomes sluggish, rinse with dilute detergent, rinse with de-ionised water and immerse the tip in a 1000ppm Iodide solution for 1 hour.

Should further conditioning be requires rub down the crystal surface with a very fine emery paper until the surface has a grey shiny appearance. Stand electrode in 100 ppm Iodide solution for 20 minutes.

Specification

Iodide is a monovalent anion. A 1 molar solution of Iodide contains 126.904 grams per litre of Iodide.

Refer to the table below for the full specification:

| Parameter | Specification |
| --- | --- |
| Overall length Body Diameter Cap Diameter Connector Cable length Resistance at 25 Deg C Concentration Range Slope Potential Drift Operating pH range Temperature range Endpoint time Interferences: Ions with coefficients above 0.001. | 155 mm 12 mm 16mm BNC 1000 mm < 2.5 Meg Ohm 0.06 to 127.000 ppm -54 to -59 mV per decade 2 mV per 8 hours 2 to 12 5 to 50 Deg C Typically 10 to 30 seconds Sulphide Cyanide |

Operating Instructions for Iodide

Ion Selective Electrode

- Introduction

  - Construction

The electrode consists of an inert polymer body on the end of which is mounted the sensitive membrane. All connections in the electrode are solid-state contacts, there being no liquids of any type inside the body. The high state of surface finish of the membrane is an essential feature for maximum sensitivity and performance. The low-noise shielded cable terminates in a BNC, though an alternative plug may be specified.

  - Operating Principles

The electrical potential of an ion selective electrode is a function of the logarithm of the activity of the ion to be measured. The relationship is given by the Nernst Equation:

RT

E(Measured)		=	Eo + zF	  in AI

where:

Eo is a constant characteristic of the electrode

R is the Universal Gas Constant

T is the Absolute Temperature

z is the charge of the ion

F is the Faraday constant

AI is the activity of the Iodide ions

This equation can be simplified to:

Em	=	Constant + S log AI

where S is the slope of the calibration curve. The measured potential Em can only be measured against a reference electrode, such as a saturated calomel electrode (SCE) placed in the same solution. Since leakage of chloride ion from the Reference Electrode will affect the chloride ions from the Reference Electrode will affect the chloride level in test solutions, a salt-bridge tube filled with 2M sodium nitrate solution must be placed between the SCE and the test solution.

Alternatively a commercial double-junction reference electrode can be used. Measurements are made with a high impedance millivoltmeter or expanded scale pH meter. A meter with a digital display is obviously more accurate and most convenient. The iodide electrode can be used in solutions containing high proportion of organic solvents without ill-effect. For further discussion on the principles of ISE measurements, the user is referred to the Technical Bulletin which is available on demand.

  - Sensitivity

The electrode will respond to uncomplexed iodide ion activity over the range 100 - 10-5M. The linear detection limit is around 5 x 10-5M. Curvature of the response is observed at lower concentrations because the solubility of the membrane material, dissolved from the membrane, contributes a small but significant amount of iodide to the solution being measured. Figure 1 shows a typical calibration plot of the electrode. Values for the slope of this plot should lie in the range 52-59mV/decade. If the total ionic-strength of all solutions is maintained constant by the addition of a concentrated (approx. 10-1M) inert electrolyte such as odium nitrate, the electrode can be calibrated directly in concentration units.

The activity coefficients of iodide ion depends upon the ionic strength of the solution. The following table can be used to calculate the approximate activity coefficients at various ionic strengths.

Ionic Strength		0.001	0.005	0.01	0.05	0.1	0.2

Activity Coeff. (f)	0.975	0.946	0.926	0.853	0.808	0.755

The ionic strength of the solution to be tested can be calculated from the formula:

Ionic Strength		=	½ Σ CiZi2

Where Ci is the concentration of species i and Zi  is its charge, and all species, cations and anions, must be taken into account.

Once the ionic strength is calculated and the activity coefficient found from Table 1, then the activity of the ion of interest is found from:

Activity (Ai)	=	Concentration (Ci) x Activity Coefficient (f)

  - Stability

The long-term drift Eo value of the iodide electrode is usually in the region of 2-4mV/day.

  - Selectivity

The only interference encountered in the use of the iodide electrode in many different situations arise from the presence of cyanide or sulphide ions. The electrode is equally sensitive to iodide and cyanide ions and so its potential will be characteristics of the total activity of both. Sulphide ion should be less than 10-7M.

Prolonged immersion of the electrode in strong > 10-2M I- solution causes a deterioration of the membrane surface. This causes a reduction in response speed. Light polishing of the membrane with very fine alumina powder on a felt pad will restore its former performance. Re-polishing can be undertaken by the manufacturers.

The electrode can be used in the pH range 2-12.

  - Response Time

The electrode response to a step change in iodide ion activity in solution is extremely rapid, with 95% of the potential change completed in 200 msec. Equilibrium is reached in less than 2 seconds. When, however, the electrode is removed from solution and placed in another of different concentration, it may take up to 20 minutes to regain equilibrium.

The presence of interfering ions does not affect the response time, but in solutions of very low ionic strength, i.e. with no ionic strength buffer, the response time may be extended to several minutes.

- Operation and Handling

  - Initial Setting-Up

On receipt of the electrode remove it from its packing and inspect for damage particularly to the highly polished membrane which forms the sensing end of the electrode. Clamp the electrode with its sensing end in a solution of sodium or potassium iodide (10-1M) for about two hours or until the electrode potential, when measured against an SCE, remains constant. Overnight soaking is recommended. After this conditioning the electrode is ready for use.

  - Making a Measurement

The electrode should first be calibrated as follows. Clamp the electrode in a vertical position such that the membrane is immersed in a beaker of 10-1M iodide solution on a magnetic stirrer. Clamp a reference electrode (a double junction Ag/AgCl or SCE may be used) so that its end is immersed in the same solution and connect both electrodes to a pH or millivolt meter. Record the potential when it becomes steady. Change the test solution sequentially for 10-2M, 10-3M and 10-4M iodide solutions and plot the calibration curve, washing and drying the electrode between solutions.

Unknown solutions can now be measured directly. For further details of methods of measurement, the user is referred to Technical Bulletin E1 available for the electrode supplier.

  - Storage and Maintenance

When the electrode is no longer needed it can be merely removed from the test solution, washed and dried and be clamped in the air or returned to its packing container. There is in principle no limit to the storage time. The membrane should be protected from scratching and abrasion during storage. The only maintenance required is an occasional polish with a tissue or soft cloth to keep the membrane surface shiny.

  - Flow-Through Configuration

A Flow-Through version is available to enable continuous flow-through measurements to be made. The reference contact can be made at any point in the solution stream joined to the flow cell by a continuous stream of liquid usually down-stream of the electrode. The combination of this electrode with a sampler and pumping systems can result in a versatile and low-cost continuous iodide monitor.

- Applications and Bibliography

The iodide electrode finds application mostly in inorganic analysis and physical measurements though some applications in biological iodide monitoring have been reported.

Inorganic Iodide Measurement

- K. Burger and B. Pinter, Hung. Sci. Instruments, 8 (1966) 11

- J. H. Woodson and H. A. Liebhafsky, Nature, 224 (1969) 690

- J. H. Woodson and H. A. Liebhafsky, Anal. Chem., 41 (1969) 1894

- M. R. Domeshwar and R. G. Dhaneshwar, Trans. Soc. Adv. Electrochem. Sci. Technol., 8(2) (1973) 61

Iodide in Non Aqueous Media

W. H. Ficklin and W. C. Gotschall, Anal. Lett., 6(3), (1973) 317

Determination of Gold with Iodide Electrode

D. Wiess, Fresenius Z. Anal. Chem., 262 (1972) 28

Determination of Sulphide with Iodide Electrode

S. Ikeda et al., Nippon Kagaku Kaishi, 8 (1973) 1473

Determination of Iodide Waters

- J. Havas et al., Hung. Sci. Inst., 9 (1967) 1923

- R. C. Harris and H. H. Williams, Appl. Meterol., 8 (1969) 299

- E. Pungar, Anal. Chem., 39 (1967) 28A

Determination of Iodide in Biological Materials

- R. A. Carter, Proc. Assoc. Clin. Biochem., 5 (1968) 67

- W. L. Hoover et al., J. Ass. Offic. Anal. Chem., 54 (1971) 760

- B. Paletta, Mikrochim Acta., 6 (1969) 1210

4. Trouble Shooting

  - Wildly erratic readings

- Air bubble trapped on outside surface of membrane

- Excessively violent stirring

  - Steady continuous drift in one direction

- Poor earthing of the mV meter

- Excessive leaking from the reference electrode junction leading to continuous changes in the activity coefficients of ions in weaker solutions.

- Temperature drift if solutions are not thermostatted, or sample vessel insulated from, for example stirrer motor.

- Not using a double-junction reference electrode or salt bridge.

  - Slope of electrode much less than 50mV/decade

- Membrane old and in need of replacing or re-polishing

- Presence of an interfering ion in a constant concentration swamping the selected ions. Check performance with pure iodide solutions.

  - Totally incomprehensible results may arise from leakage of solution into electrode via the membrane/body seal. If all fails to cure a problem please ship the electrode back to the supplier for exchange.

# Text boxes

150mV

300mV

Mono/Reference

Combination

SENTEK

Unit 6 & 7 Crittall Court, Crittall Drive, Springwood Industrial

Estate, Braintree, Essex, CM7 2SE

Tel: +44 (0) 1376 340 456  Fax: +44 (0) 1376 340 453  Email:  sentekuk@btconnect.com

SENTEK

Unit 6 & 7 Crittall Court, Crittall Drive, Springwood Industrial

Estate, Braintree, Essex, CM7 2SE

Tel: + 44 (0) 1376 340 456  Fax:  +44 (0) 1376 340 453  Email: sentekuk@btconnect.com